## Slide 1: DDS Debugging
AD9910 Verilog & Board Debug

## Slide 2: Glitch Problem
Profile 000->011 makes 001 010 oftenly

## Slide 3: Possible Causes
1. Metastable → need SYNC utility of AD9910 to solve
2. Transmission line → low possibility
3. FPGA problem → ARTY S7 connected to DDS had problem

## Slide 4: Metastable
Simplified flip-flop model
S1 close, S2 open works as a clock edge event
D flip flop
D Q
CLK
CLK
D
Q

## Slide 5: Metastable
D flip flop
D Q
CLK
CLK
D
Q
setup
hold
Input Signal should not change in setup hold time
→setup hold time condition

## Slide 6: Final Cause
Different time of profile signal arrival time
→ Only one profile pin cause change in input buffer register
see https://ez.analog.com/dds/f/q-a/29655/ad9914-profile-change-problem
NIST and M-Labs suffered from same issue with AD9914 and AD9910 respectively. see https://github.com/m-labs/artiq/issues/1994
Profile[2]
register
Profile[2]
Profile[1]
Profile[0]
SYNC_CLK
Profile[1]
register
setup
hold
setup-hold time violation!
metastable
if setup+hold ~= 2ns -> 1-(8/10)*(8/10) = 9/25 -> 36% of glitch probability

## Slide 7: Non perfect solution
Send IO Update Manually
Do not use internal IO Update utility!(make CFR2[23] = 0)
Same with using two D flip-flop

## Slide 8: Experiment for timing
Validation of our assumption
1GHz CLOCK
delay control
AD9910
ARTYS7
spectrum analyzer
profile change
000
010
001
011
profile
f1
f2
f3
f4
frequency
adjust delay

## Slide 9: Delay Control Circuit for Experiment(1)
Is it possible to make without PCB?
Is inductance of jumper acceptable?

## Slide 10: Delay Control Circuit for Experiment(2)
Breadboard inductance, capacitance does not matter?
Working frequency is to low!

## Slide 11: Delay Control Circuit for Experiment(3)
Use SG384(clock source) phase adjustment
Simple solution!

## Slide 12: Experiment for Solution Validation
Validation of our assumption
1GHz CLOCK
AD9910
ARTYS7
spectrum analyzer
000
010
001
011
profile
f1
f2
f3
f4
frequency
D flip-flop

## Slide 13: Solution1
ARTY S7
AD9910
AD9910
OSERDES
ISERDES
PROFILE
PROFILE
SYNC_CLK
SYNC_CLK
SYNC_OUT
SYNC_OUT
SYNC_IN
SYNC_IN
1ns control
250MHz
62.5MHz
No SERDES in Spartan7
Impossible!

## Slide 14: Solution2
ARTY S7
AD9910
AD9910
PLL
IOUPDATE
PROFILE
IOUPDATE
SYNC_OUT
SYNC_OUT
SYNC_IN
SYNC_IN
62.5MHz
62.5MHz
Does it work properly? (PLL jitter of FPGA is sufficiently low?)
If this makes noise to AC signal, is it worth that much?
control SYNC_IN delay

## Slide 15: Solution3
ARTY S7
AD9910
AD9910
IOUPDATE
PROFILE
SYNC_OUT
SYNC_OUT
SYNC_IN
SYNC_IN
62.5MHz
New PCB should be made
No proper delay controller exist to control each signal(too expensive)-> should be made in trace
buffer
CPLD
PMOD
FMC
Digital
Power
IOUPDATE
PROFILE
delay control
SPI
SPI
Oscillator
Pmod Cable
FMC Cable
e.g.DS1023, MC100EP195B
or trace delay

## Slide 16: Solution4
ARTY S7
AD9910
AD9910
IOUPDATE
PROFILE
SYNC_OUT
SYNC_OUT
SYNC_IN
SYNC_IN
62.5MHz
New PCB should be made
Only one proper timing delay chip is required!
However we cannot use this solution if clock phase of FPGA and DDS board is not synchronized well…
CPLD
Digital
Power
IOUPDATE
PROFILE
delay control
SPI
SPI
Oscillator
CLOCK SOURCE
SYSCLK

## Slide 17: Solution5
ARTY S7
AD9910
AD9910
IOUPDATE
PROFILE
SYNC_OUT
SYNC_OUT
SYNC_CLK
SYNC_IN
62.5MHz
It just lower metastable probability, does not fix it..
But this seems to be sufficient in this problem!
https://www.mouser.kr/ProductDetail/Nexperia/74AUP1G79GV125?qs=jquClx72t9COASqZ9WZ81A%3D%3D
IOUPDATE
PROFILE
SPI
SPI
CPLD
Digital
Power
D flip-flop
(74AUP1G79)
Oscillator

## Slide 18: More improvement
NEW FPGA
AD9910
AD9910
IOUPDATE
PROFILE
IOUPDATE
PROFILE
SYNC_OUT
SYNC_OUT
SYNC_IN
SYNC_IN
OUTPUT
OUTPUT
LPF
Attenuator
Amplifier
AD9910
IOUPDATE
PROFILE
SYNC_OUT
SYNC_IN
OUTPUT
AD9910
IOUPDATE
PROFILE
SYNC_OUT
SYNC_IN
OUTPUT
External
Clock
Oscillator
SMA
CPLD
CLK_DELAY
CLK_DELAY
CLK_DELAY
CLK_DELAY
CLK_DELAY
CLK_DELAY
CLK_DELAY
REF_CLK
REF_CLK
REF_CLK
REF_CLK
CLK
DIVIDER
EXT
CLK

## Slide 19: Clock delay chip
AD9515 -> 1GHz output does not work with delay adjustment
HMC856LC5 -> Nice but Too expensive!(~$300)

## Slide 20: Clock delay chip
Chip | frequency [GHz] | delay step [ps] | total delay [ns] | jitter [ps] | INPUT | OUTPUT | Price [Won] | 
SY100EP195VTG | 1.6 | 10 | 2.1~10.8 | 0.2~1 | PECL, LVPECL, NECL, LVNECL | ECL | 20,000 | cont ctrl
SY100EP196VTG | 2.5 | 10 | 2.2~12.2 | 0.2~1 | PECL, LVPECL, NECL, LVNECL | PECL, NECL | 24,000 | cont ctrl
SY100E196JY | 1 | 20 | 1.2~3.3 | 5 | ECL | ECL? | 10,000 | cont ctrl
NB6L295 | 1.5 | 11 | 6~17.2 | 3 | LVPECL, CMS, LVDS | LVPECL, LVCMOS, LVTTL | 40,000 | SPI interface
MC100EP195 | 1.2 | 10 | 2.2~12.4 | 1.2 | LVCMOS, ECL, LVTTL | ECL? | 34,000 | 
MC100EP195B | 1.2 | 10 | 2.2 | 1.1~5 | LVPECL, LVNECL, LVDS | ECL | 39,000 | 
MC100EP196 | 1.2 | 10 | 2.2~12.2 | 3 | NECL, PECL | ECL | 23,000 | cont ctrl
MC100EP196B | 1.2 | 10 | 2.2~12.4 | 1.1~5 | NECL, PECL | ECL | 29,000 | cont ctrl
SY89295U | 1.5 | 10 | 3.3~14.8 | 2~10 | LVPECL | LVPECL | 20,000 | 
SY89296U | 1.6 | 10 | 3.2~14.8 | 2~10 | LVPECL | LVPECL | 22,000 | 
SY89297UMH | 1.6 | 5 | 2~7 | 1 | CML, LVPECL, LVDS | CML | 33,000 | SPI interface

## Slide 21: Clock delay chip
Chip | frequency [GHz] | delay step [ps] | total delay [ns] | jitter [ps] | INPUT | OUTPUT | Price [Won] | 
SY100EP195VTG | 1.6 | 10 | 2.1~10.8 | 0.2~1 | PECL, LVPECL, NECL, LVNECL | ECL | 20,000 | cont ctrl
SY100EP196VTG | 2.5 | 10 | 2.2~12.2 | 0.2~1 | PECL, LVPECL, NECL, LVNECL | PECL, NECL | 24,000 | cont ctrl
SY100E196JY | 1 | 20 | 1.2~3.3 | 5 | ECL | ECL? | 10,000 | cont ctrl
NB6L295 | 1.5 | 11 | 6~17.2 | 3 | LVPECL, CMS, LVDS | LVPECL, LVCMOS, LVTTL | 40,000 | SPI interface
MC100EP195 | 1.2 | 10 | 2.2~12.4 | 1.2 | LVCMOS, ECL, LVTTL | ECL? | 34,000 | 
MC100EP195B | 1.2 | 10 | 2.2 | 1.1~5 | LVPECL, LVNECL, LVDS | ECL | 39,000 | 
MC100EP196 | 1.2 | 10 | 2.2~12.2 | 3 | NECL, PECL | ECL | 23,000 | cont ctrl
MC100EP196B | 1.2 | 10 | 2.2~12.4 | 1.1~5 | NECL, PECL | ECL | 29,000 | cont ctrl
SY89295U | 1.5 | 10 | 3.3~14.8 | 2~10 | LVPECL | LVPECL | 20,000 | 
SY89296U | 1.6 | 10 | 3.2~14.8 | 2~10 | LVPECL | LVPECL | 22,000 | 
SY89297UMH | 1.6 | 5 | 2~7 | 1 | CML, LVPECL, LVDS | CML | 33,000 | SPI interface
cont ctrl require additional circuit(e.g. DAC)

## Slide 22: Voltage range

## Slide 23: MC100EP195
D0~D10, LEN, EN, CASCADE*2 ->15
RF -> 1 (using invertor) or2
MAX : 15*4 + 2*5= 70 ( for 4 CLK)
-> extending to 7 CLK delay require more than 100 pins

## Slide 24: NB6L295
SDATA, SCLK, SLOAD, EN ->4
RF -> 1 (using invertor) or2
MAX : 4*4 + 2*5= 26 ( for 4 CLK)
-> extending to 7 CLK delay requires almost 28 pins

## Slide 25: Circuit(1)
jb_1 | jb_5
jb_2 | jb_6
jb_3 | jb_7
jb_4 | jb_8
GND | GND
VCC | VCC
CPLD
(XC9536XL or XC2C32A
or XC2C64A or XC2C128A)
5
1
6
2
7
3
8
4
G
G
V
V
delay chip
digital
analog
voltage regulator
JTAG
SPDT
bypass capacitor
bypass capacitor
shield
ground is separated
sma
Ntype
sma
Ntype
oscillator
-3dBm
-3dBm
Impedance matched at all port
Total -6dBm loss
voltage regulator
power is separated
decoupling capacitor
decoupling capacitor

## Slide 26: Wilkinson divider -> RF SPDT switch
-3dBm
HMC545A
QPC7522
HSWA2-30DR+
Low insertion loss
Digital signal in analog section increase
Does RF SPDT makes random jitter?

## Slide 27: Circuit(2)
jb_1 | jb_5
jb_2 | jb_6
jb_3 | jb_7
jb_4 | jb_8
GND | GND
VCC | VCC
XC2C128A
5
1
6
2
7
3
8
4
G
G
V
V
NB6L295
digital
analog
voltage regulator
JTAG
SPDT
bypass capacitor
bypass capacitor
shield
ground is separated
sma
Ntype
sma
oscillator
-3dBm
Total -6dBm loss
voltage regulator
power is separated
decoupling capacitor
decoupling capacitor
NB6L295
NB6L295
NB6L295
RF
SPDT

## Slide 28: Circuit(3)
jb_1 | jb_5
jb_2 | jb_6
jb_3 | jb_7
jb_4 | jb_8
GND | GND
VCC | VCC
XC2C128A
5
1
6
2
7
3
8
4
G
G
V
V
NB6L295
digital
analog
voltage regulator
JTAG
bypass capacitor
bypass capacitor
shield
ground is separated
sma
sma
oscillator
-3dBm
voltage regulator
power is separated
decoupling capacitor
decoupling capacitor
NB6L295
NB6L295
NB6L295
DC
Impedance matched at all port
Total -6dBm loss

## Slide 29: RF SPDT switch
use 6 layer PCB
L4 layer
RFc
RF1
RF2
GND
B
A
Top Layer
for equal timing length is lengthened

## Slide 30: Program
PCB → Kicad 6.0
CPLD → ISE14.7 in Ubuntu 22.04.1LTS
FPGA → Vivado2017.3

## Slide 31
step down switching
temperature sensor & thermal watchdog
step down switch
low dropout linear regulator
step down switch
ultrahigh PSRR linear regulator
micropower
regulator

## Slide 32: Regulator

## Slide 33
Bypass capacitor
Noise suprress inductor
decoupling capacitor

## Slide 35: Synchronization of AD9910
SYNC pin should be used
SYNC_IN
SYNC_SMP_ERROR

## Slide 36: Urukul Board
Not seems to be considering RF effect that much

## Slide 37: AD9910 Block diagram
Output
FUD Pin
Data I/O pin
Register select pin
OSK enable pin
Clock input
Synchronization pin
(SYNC_IN != SYNC_CLK)
Not used
DRG control pin

## Slide 38: I/O update & SYNC_CLK timing
DDS1 sense I/O_UPDATE
DDS2 sense I/O_UPDATE
DDS1 update register
DDS2 update register
SYS_CLK
(Same in DDS1,2)
DDS1 SYNC_CLK
DDS2 SYNC_CLK
Common I/O update
SYNC_CLK = SYS_CLK/4
SYNC_OUT = SYS_CLK/16 (Controllable)
Even with phase aligned SYS_CLK, synchornization can fail

## Slide 39: Synchronization of AD9910
Check whether hold/setup
condition satisfies
Rising edge of SYNC_IN starts counter→ make
coincidence SYNC_CLK
SYNC RECEIVER DELAY
SYNC_CLK
when SYNC RECEIVER DELAY != 0
SYNC_CLK
when SYNC RECEIVER DELAY = 0
SYS_CLK
Control SYNC_IN delay
Control skew of SYNC_IN, SYS_CLK
metastable can occur
see https://ez.analog.com/dds/f/q-a/30165/ad9957-multichip-sync

## Slide 40: SYNC Validation
Skew caused by SYNC_VALIDATION_DELAY
SYS_CLK
SYNC VALIDATION DELAY
SYNC VALIDATION DELAY
SYS_CLK
SYNC_IN
SYNC_IN
SETUP VALIDATION
HOLD VALIDATION

## Slide 41: Synchronizing
Case1
Case2
Sync Validation delay
0
1
2
3
4
5
SYNC_IN delay
0 1 2 3 4 5 6 7 8 9 10 11 12 13 14 15 16 17 18 19 20 21 22 23 24 25 26 27 28 29 30 31
Sync Validation delay
0
1
2
3
4
5
SYNC_IN delay
0 1 2 3 4 5 6 7 8 9 10 11 12 13 14 15 16 17 18 19 20 21 22 23 24 25 26 27 28 29 30 31
SYNC_SMP_ERR

## Slide 42: tune_sync_delay function
759 @kernel
760 def tune_sync_delay(self, search_seed=15):
776 if not self.cpld.sync_div:
777 raise ValueError("parent cpld does not drive SYNC")
778 search_span = 31
781 min_window = 0
782 margin = 1
783 for window in range(16):
784 next_seed = -1
785 for in_delay in range(search_span - 2*window):
787 if in_delay & 1:
788 in_delay = -in_delay
789 in_delay = search_seed + (in_delay >> 1)
790 if in_delay < 0 or in_delay > 31:
791 continue
792 self.set_sync(in_delay, window)
793 self.clear_smp_err()
795 delay(100*us)
796 err = urukul_sta_smp_err(self.cpld.sta_read())
797 delay(100*us) # slack
798 if not (err >> (self.chip_select - 4)) & 1:
799 next_seed = in_delay
800 break
801 if next_seed >= 0:
802 search_seed = next_seed
803 continue
804 elif window > min_window:
806 window = max(min_window, window - 1 - margin)
807 self.set_sync(search_seed, window)
808 self.clear_smp_err()
809 delay(100*us)
810 return search_seed, window
811 else:
812 break
813 raise ValueError("no valid window/delay")

## Slide 43: I/O_UPDATE delay
Frequency can be odd or even
Even number of SYNC_CLK
Odd Frequency
Odd Frequency
Even Frequency
1FTW
SYS_CLK
DDS SYNC_CLK
Common I/O update
FTW

## Slide 44: I/O_UPDATE delay
Frequency can be only even
Even number of SYNC_CLK
Odd Frequency
Even Frequency
1FTW
I/O_UPDATE DELAY
SYS_CLK
DDS SYNC_CLK
Common I/O update
FTW

## Slide 45: tune_io_update_delay function
882 @kernel
883 def tune_io_update_delay(self):
899 period = self.sysclk_per_mu * 4
900 repeat = 100
901 for i in range(period):
902 t = 0
904 for j in range(repeat):
905 t += self.measure_io_update_alignment(i, i + 2)
906 if t != 0:
907 continue
909 t1 = [0, 0]
910 for j in range(repeat):
911 t1[0] += self.measure_io_update_alignment(i, i + 1)
912 t1[1] += self.measure_io_update_alignment(i + 1, i + 2)
913 if ((t1[0] == 0 and t1[1] == 0) or
914 (t1[0] == repeat and t1[1] == repeat)):
916 raise ValueError("no clear IO_UPDATE-SYNC_CLK alignment edge found")
918 else:
920 return (i + 1 + period//2) & (period - 1)
921 raise ValueError("no IO_UPDATE-SYNC_CLK alignment edge found")

## Slide 46: measure_io_update_alignment function
815 @kernel
816 def measure_io_update_alignment(self, delay_start, delay_stop):
830 self.set_cfr1(drg_load_lrr=1, drg_autoclear=1)
832 self.write32(_AD9910_REG_CFR2, 0x01090000)
859 self.write64(_AD9910_REG_RAMP_LIMIT, -1, 0)
861 self.write32(_AD9910_REG_RAMP_RATE, 0x00010000)
863 self.write64(_AD9910_REG_RAMP_STEP, -1, 0)
865 t = now_mu() + 8 & ~7
866 at_mu(t + delay_start)
868 self.cpld.io_update.pulse_mu(16 - delay_start)
870 self.set_cfr1()
872 self.write64(_AD9910_REG_RAMP_STEP, 0, 0)
873 at_mu(t + 0x1000 + delay_stop)
874 self.cpld.io_update.pulse_mu(16 - delay_stop)
875 ftw = self.read32(_AD9910_REG_FTW)
876 delay(100*us)
878 self.write32(_AD9910_REG_CFR2, 0x01010000)
879 self.cpld.io_update.pulse_mu(8)
880 return ftw & 1

## Slide 47: Urukul:AD9910 Parameter Control(1)
1 from artiq.experiment import *
2
3
4 class DDSTest(EnvExperiment):
5 def build(self):
6 self.setattr_device("core")
7 self.setattr_device("urukul0_cpld")
8 self.urukul_dds0 = self.get_device("urukul0_ch0")
9
10 @kernel
11 def run(self):
14 self.urukul0_cpld.init()
15
16 delay(1*ms)
17 self.urukul_dds0.init()
18
19 self.urukul_dds0.set(frequency = 100*MHz, phase = 0.0, profile = 0)
20 self.urukul_dds0.set(frequency = 150*MHz, phase = 0.0, profile = 1)
21 self.urukul_dds0.set(frequency = 200*MHz, phase = 0.0, profile = 2)
22 self.urukul_dds0.set(frequency = 200*MHz, phase = 0.5, profile = 3)
23 self.urukul_dds0.sw.on()
24
profile | frequency | phase | amplitude
0 | 100MHz | 0 | 1.0
1 | 150MHz | 0 | 1.0
2 | 200MHz | 0 | 1.0
3 | 200MHz | pi | 1.0

## Slide 48: Urukul:AD9910 Parameter Control(2)
25 delay(0.01*us)
26 self.urukul0_cpld.set_profile(profile = 1)
27 self.urukul_dds0.set_att(6.)
28
29 delay(0.01*us)
30 self.urukul0_cpld.set_profile(profile = 2)
31
32 delay(0.01*us)
33 self.urukul0_cpld.set_profile(profile = 3)
34
35 delay(0.01*us)
36 self.urukul_dds0.sw.off()
set which profile register to supply parameter
profile | frequency | phase | amplitude
0 | 100MHz | 0 | 1.0
1 | 150MHz | 0 | 1.0
2 | 200MHz | 0 | 1.0
3 | 200MHz | pi | 1.0

## Slide 49: Urukul:AD9910 Parameter Control(2)
25 delay(0.01*us)
26 self.urukul0_cpld.set_profile(profile = 1)
27 self.urukul_dds0.set_att(6.)
28
29 delay(0.01*us)
30 self.urukul0_cpld.set_profile(profile = 2)
31
32 delay(0.01*us)
33 self.urukul0_cpld.set_profile(profile = 3)
34
35 delay(0.01*us)
36 self.urukul_dds0.sw.off()
set which profile register to supply parameter
profile | frequency | phase | amplitude
0 | 100MHz | 0 | 1.0
1 | 150MHz | 0 | 1.0
2 | 200MHz | 0 | 1.0
3 | 200MHz | pi | 1.0

## Slide 50: Urukul:AD9910 Parameter Control(2)
25 delay(0.01*us)
26 self.urukul0_cpld.set_profile(profile = 1)
27 self.urukul_dds0.set_att(6.)
28
29 delay(0.01*us)
30 self.urukul0_cpld.set_profile(profile = 2)
31
32 delay(0.01*us)
33 self.urukul0_cpld.set_profile(profile = 3)
34
35 delay(0.01*us)
36 self.urukul_dds0.sw.off()
set which profile register to supply parameter
profile | frequency | phase | amplitude
0 | 100MHz | 0 | 1.0
1 | 150MHz | 0 | 1.0
2 | 200MHz | 0 | 1.0
3 | 200MHz | pi | 1.0

## Slide 51: Urukul:AD9910 RAM Mode(1)
1 from artiq.experiment import *
2
3
4 class DDSTest(EnvExperiment):
5 def build(self):
6 self.setattr_device("core")
7 self.setattr_device("urukul0_cpld")
8 self.urukul_dds0 = self.get_device("urukul0_ch0")
9
10 @kernel
11 def run(self):
14 self.urukul0_cpld.init()
15 delay(1*ms)
16
17 self.urukul_dds0.init()
18 delay(1*us)
19
20 self.urukul_dds0.set_profile_ram(start = 0, end = 3, step = 2, profile = 0, nodwell_high = 0, zero_crossing = 0, mode = 1)
21 delay(1*us)
22
23 self.urukul_dds0.set_profile_ram(start = 4, end = 7, step = 1, profile = 1, nodwell_high = 0, zero_crossing = 0, mode = 1)
24 delay(1*us)
25
26 self.urukul_dds0.set_profile_ram(start = 8, end = 11, step = 2, profile = 2, nodwell_high = 0, zero_crossing = 0, mode = 1)
27
set profile for ram mode

## Slide 52: Urukul:AD9910 RAM Mode(2)
28 ram_data_SI = [
29 50*MHz, 60*MHz, 70*MHz, 80*MHz, 90*MHz, 100*MHz, 110*MHz, 120*MHz, 130*MHz, 140*MHz, 150*MHz, 160*MHz
30 ]
31 ram_data_MU = list()
32
33 self.urukul_dds0.frequency_to_ram(ram_data_SI, ram_data_MU)
34 self.urukul_dds0.write_ram(ram_data_MU)
35
36 self.urukul_dds0.set_asf( self.urukul_dds0.amplitude_to_asf( amplitude = 1.0 ) )
37 self.urukul_dds0.set_pow( self.urukul_dds0.turns_to_pow( turns = 0.0 ) )
38
39 self.urukul_dds0.set_cfr1(power_down = 0b0000,
40 phase_autoclear = 0,
41 drg_load_lrr = 0,
42 drg_autoclear = 0,
43 internal_profile = 0,
44 ram_destination = 0,
45 ram_enable = 1,
46 )
47
48 delay(0.1*us)
49
50 self.urukul_dds0.sw.on()
51 self.cpld.io_update.pulse(1*us)
frequency to save in RAM with SI unit
frequency to save in RAM with machine unit
save amplitude, phase to supply
in RAM mode
set Control Function Register
to RAM mode
IO update
RAM | DATA
0 | 50MHz
1 | 60MHz
2 | 70MHz
3 | 80MHz
4 | 90MHz
5 | 100MHz
6 | 110MHz
7 | 120MHz
8 | 130MHz
9 | 140MHz
10 | 150MHz
11 | 160MHz

## Slide 53: Urukul:AD9910 Synchronization Code(1)
1 from artiq.experiment import *
2
3 _AD9910_REG_PROFILE0 = 0x0e
4 _AD9910_REG_PROFILE1 = 0x0f
5 _AD9910_REG_PROFILE2 = 0x10
6 _AD9910_REG_PROFILE3 = 0x11
7 _AD9910_REG_PROFILE4 = 0x12
8 _AD9910_REG_PROFILE5 = 0x13
9 _AD9910_REG_PROFILE6 = 0x14
10 _AD9910_REG_PROFILE7 = 0x15
11
12
13 class DDSTest(EnvExperiment):
14 def build(self):
15 self.setattr_device("core")
16 self.setattr_device("urukul0_cpld")
17 self.dds_set = [ self.get_device("urukul0_ch" + str(i)) for i in range(4) ]
18
19 @kernel
20 def run(self):
23 self.urukul0_cpld.init()
24 delay(1*ms)
25
26 for dds in self.dds_set:
27 dds.init()
28 delay(1*ms)
29

## Slide 54: Urukul:AD9910 Synchronization Code(1)
1 from artiq.experiment import *
2
3 _AD9910_REG_PROFILE0 = 0x0e
4 _AD9910_REG_PROFILE1 = 0x0f
5 _AD9910_REG_PROFILE2 = 0x10
6 _AD9910_REG_PROFILE3 = 0x11
7 _AD9910_REG_PROFILE4 = 0x12
8 _AD9910_REG_PROFILE5 = 0x13
9 _AD9910_REG_PROFILE6 = 0x14
10 _AD9910_REG_PROFILE7 = 0x15
11
12
13 class DDSTest(EnvExperiment):
14 def build(self):
15 self.setattr_device("core")
16 self.setattr_device("urukul0_cpld")
17 self.dds_set = [ self.get_device("urukul0_ch" + str(i)) for i in range(4) ]
18
19 @kernel
20 def run(self):
23 self.urukul0_cpld.init()
24 delay(1*ms)
25
26 for dds in self.dds_set:
27 dds.init()
28 delay(1*ms)
29

## Slide 55: Urukul:AD9910 Synchronization Code(2)
30 temp_sum = 0
31 amplitude = [0.4, 0.6, 0.8, 1.0]
32 freq = [50*MHz, 80*MHz, 110*MHz, 140*MHz]
33 phase = [0.0, 0.2, 0.4, 0.6]
34 attenuation = [ 6.0, 12.0, 18.0, 24.0 ]
35
36 for i in range(4):
37 temp_sync_delay, temp_validation_delay = self.dds_set[i].tune_sync_delay()
38 delay(1*ms)
39 temp_io_update_delay = self.dds_set[i].tune_io_update_delay()
40 self.dds_set[i].sync_data.sync_delay_seed = temp_sync_delay
41 self.dds_set[i].sync_data.io_update_delay = temp_validation_delay
42 temp_sum += temp_io_update_delay
43 temp_asf = self.dds_set[i].amplitude_to_asf(amplitude[i])
44 temp_ftw = self.dds_set[i].frequency_to_ftw(freq[i])
45 temp_pow = self.dds_set[i].turns_to_pow(phase[i])
46 profile_num = _AD9910_REG_PROFILE0 + i
47 self.dds_set[i].write64( profile_num, ( ( temp_asf & 0x3fff ) << 16 ) | ( temp_pow & 0xffff ), temp_ftw & 0xffffffff )
48 delay(1*us)
49 self.dds_set[i].set_att(attenuation[i])
50 delay(1*us)
51
variable to calculate average IO_UPDATE signal delay
parameters to set DDS
find and set proper SYNC_IN delay for each DDS
find and set proper IO_UPDATE delay for each DDS

## Slide 56: Urukul:AD9910 Synchronization Code(3)
52 io_update_delay_avg = int64(round(temp_sum/4))
53
54 t = now_mu() + int64(8) + io_update_delay_avg
55 at_mu(t)
56 self.urukul0_cpld.io_update.pulse_mu(8)
57
58 for dds in self.dds_set:
59 dds.sw.on()

## Slide 57: Urukul DeviceDB(1)
1 core_addr = "192.168.1.70"
2
3 device_db = {
4 "core": {
5 "type": "local",
6 "module": "artiq.coredevice.core",
7 "class": "Core",
8 "arguments": {"host": core_addr, "ref_period": 1/(8*150e6)}
9 },
10 "core_log": {
11 "type": "controller",
12 "host": "::1",
13 "port": 1068,
14 "command": "aqctl_corelog -p {port} --bind {bind} " + core_addr
15 },
16 "core_cache": {
17 "type": "local",
18 "module": "artiq.coredevice.cache",
19 "class": "CoreCache"
20 },
21 "core_dma": {
22 "type": "local",
23 "module": "artiq.coredevice.dma",
24 "class": "CoreDMA"
25 },
26 }
27
module where class is declared
arguments for class
-host : host IP
-ref_period : clock period of RTIO
class to reference

## Slide 58: Urukul DeviceDB(2)
28 for i in range(8):
29 device_db["ttl" + str(i)] = {
30 "type": "local",
31 "module": "artiq.coredevice.ttl",
32 "class": "TTLInOut" if i < 4 else "TTLOut",
33 "arguments": {"channel": 1+i},
34 }
35
36 device_db.update(
37 spi_urukul0={
38 "type": "local",
39 "module": "artiq.coredevice.spi2",
40 "class": "SPIMaster",
41 "arguments": {"channel": 9}
42 },
43 ttl_urukul0_io_update={
44 "type": "local",
45 "module": "artiq.coredevice.ttl",
46 "class": "TTLOut",
47 "arguments": {"channel": 10}
48 },
49 ttl_urukul0_sw0={
50 "type": "local",
51 "module": "artiq.coredevice.ttl",
52 "class": "TTLOut",
53 "arguments": {"channel": 11}
54 },
IO_UPDATE pin
SPI interface pin (4-wire)
RF switch pin

## Slide 59: DeviceDB(3)
55 ttl_urukul0_sw1={
56 "type": "local",
57 "module": "artiq.coredevice.ttl",
58 "class": "TTLOut",
59 "arguments": {"channel": 12}
60 },
61 ttl_urukul0_sw2={
62 "type": "local",
63 "module": "artiq.coredevice.ttl",
64 "class": "TTLOut",
65 "arguments": {"channel": 13}
66 },
67 ttl_urukul0_sw3={
68 "type": "local",
69 "module": "artiq.coredevice.ttl",
70 "class": "TTLOut",
71 "arguments": {"channel": 14}
72 },
73 urukul0_cpld={
74 "type": "local",
75 "module": "artiq.coredevice.urukul",
76 "class": "CPLD",
77 "arguments": {
78 "spi_device": "spi_urukul0",
79 "io_update_device": "ttl_urukul0_io_update",
80 "refclk": 150e6,
81 "clk_sel": 2
82 }
83 }
84 )
SPI bus for CPLD
IO update channel
select clock source
-0 : Onboard 100MHz Osillator
-1 : External SMA
-2 : Internal MMCX
AD9910 chip
CPLD chip (XC2C256-6FTG256C)
Attenuator chip (HMC542BLP4E)
Preamp (ERA-3SM+)
Internal Clock (MMCX) -> FPGA
External Clock (SMA)
Onboard Oscillator (CCHD950)
Low pass filter (LFCN-320 or User Defined)
RF Switch (HMC349ALP4CE)

## Slide 60: DeviceDB(3)
86 device_db["urukul0_ch0"] = {
87 "type": "local",
88 "module": "artiq.coredevice.ad9912",
89 "class": "AD9912",
90 "arguments": {
91 "pll_n": 4,
92 "chip_select": 4,
93 "cpld_device": "urukul0_cpld",
94 "sw_device": "ttl_urukul0_sw0",
95 }
96 }
97
98 device_db["urukul0_ch1"] = {
99 "type": "local",
100 "module": "artiq.coredevice.ad9912",
101 "class": "AD9912",
102 "arguments": {
103 "pll_n": 4,
104 "chip_select": 5,
105 "cpld_device": "urukul0_cpld",
106 "sw_device": "ttl_urukul0_sw0",
107 }
108 }
109
110 device_db["urukul0_ch2"] = {
111 "type": "local",
112 "module": "artiq.coredevice.ad9912",
113 "class": "AD9912",
114 "arguments": {
115 "pll_n": 4,
116 "chip_select": 6,
117 "cpld_device": "urukul0_cpld",
118 "sw_device": "ttl_urukul0_sw0",
119 }
120 }
121
122 device_db["urukul0_ch3"] = {
123 "type": "local",
124 "module": "artiq.coredevice.ad9912",
125 "class": "AD9912",
126 "arguments": {
127 "pll_n": 4,
128 "chip_select": 7,
129 "cpld_device": "urukul0_cpld",
130 "sw_device": "ttl_urukul0_sw0",
131 }
132 }
RF switch driver
CPLD driver
PLL multiply factor

## Slide 61: AD9912 constructor
11 class AD9912:
30 def __init__(self, dmgr, chip_select, cpld_device, sw_device=None, pll_n=10):
31
32 self.cpld = dmgr.get(cpld_device)
33 self.core = self.cpld.core
34 self.bus = self.cpld.bus
35 assert 4 <= chip_select <= 7
36 self.chip_select = chip_select
37 if sw_device:
38 self.sw = dmgr.get(sw_device)
39 self.kernel_invariants.add("sw")
40 self.pll_n = pll_n
41 sysclk = self.cpld.refclk/[1, 1, 2, 4][self.cpld.clk_div]*pll_n
42 assert sysclk <= 1e9
43 self.ftw_per_hz = 1/sysclk*(int64(1) << 48)
44
set SYSCLK of DDS -> in this case SYSCLK = 600MHz

## Slide 62: Urukul:AD9912 Parameter Control(1)
1 from artiq.experiment import *
2
3
4 class DDSTest(EnvExperiment):
5 def build(self):
6 self.setattr_device("core")
7 self.setattr_device("urukul0_cpld")
8 self.urukul_dds0 = self.get_device("urukul0_ch0")
9 self.urukul_dds1 = self.get_device("urukul0_ch1")
10 self.urukul_dds2 = self.get_device("urukul0_ch2")
11 self.urukul_dds3 = self.get_device("urukul0_ch3")
12
13 @kernel
14 def run(self):
17 self.urukul0_cpld.init()
18
19 delay(1*ms)
20 self.urukul_dds0.init()
21 delay(1*ms)
22 self.urukul_dds1.init()
23 delay(1*ms)
24 self.urukul_dds2.init()
25 delay(1*ms)
26 self.urukul_dds3.init()
27 delay(1*ms)

## Slide 63: Urukul:AD9912 Parameter Control(2)
29 self.urukul_dds0.set(frequency = 100*MHz, phase = 0.0)
30 self.urukul_dds0.sw.on()
31 delay(0.1*us)
32 #self.urukul_dds0.cfg_sw(1)
33 self.urukul_dds1.set(frequency = 100*MHz, phase = 0.5)
34 self.urukul_dds1.sw.on()
35 delay(0.1*us)
36 self.urukul_dds2.set(frequency = 200*MHz, phase = 0.0)
37 self.urukul_dds2.sw.on()
38 delay(0.1*us)
39 self.urukul_dds3.set(frequency = 200*MHz, phase = 0.75)
40 self.urukul_dds3.sw.on()
41 delay(0.1*us)
42
43 self.urukul_dds0.set_att(6.0)
44 self.urukul_dds1.set_att(12.0)
45 self.urukul_dds2.set_att(18.0)
46 self.urukul_dds3.set_att(24.0)
balun
-0.5dB
LPF
-0.5dB
T-pad
-9dB
Amp
+23dB
(P1dB +13dBm)
attenuator
-0~31.5dB
CPLD
chip
DDS chip
SMA output
change CPLD register value
-actual gain = +23 – 0.5 – 0.5 – 6 – 9 = +7dB

## Slide 64: Urukul:AD9912 Parameter Control(2)
balun
-0.5dB
LPF
-0.5dB
T-pad
-9dB
Amp
+23dB
(P1dB +13dBm)
attenuator
-0~31.5dB
CPLD
chip
DDS chip
SMA output
29 self.urukul_dds0.set(frequency = 100*MHz, phase = 0.0)
30 self.urukul_dds0.sw.on()
31 delay(0.1*us)
32 #self.urukul_dds0.cfg_sw(1)
33 self.urukul_dds1.set(frequency = 100*MHz, phase = 0.5)
34 self.urukul_dds1.sw.on()
35 delay(0.1*us)
36 self.urukul_dds2.set(frequency = 200*MHz, phase = 0.0)
37 self.urukul_dds2.sw.on()
38 delay(0.1*us)
39 self.urukul_dds3.set(frequency = 200*MHz, phase = 0.75)
40 self.urukul_dds3.sw.on()
41 delay(0.1*us)
42
43 self.urukul_dds0.set_att(6.0)
44 self.urukul_dds1.set_att(12.0)
45 self.urukul_dds2.set_att(18.0)
46 self.urukul_dds3.set_att(24.0)
change CPLD register value
-actual gain = +23 – 0.5 – 0.5 – 6 – 9 = +7dB

## Slide 65: Urukul Output Channel Circuit

## Slide 66: Things to be considered
Phase of RF signal to AOM is important…?
200MHz -> l = 1.5m -> l/6= 25cm => RF effect considerable?
Frequency value should be written in integer(float number is not exact)
Vivado version 2019.1 or 2017.3?
Clock on board
LPF Chip
Attenuator with Amplifier
Connector on board

## Slide 67: Things doing
ILA on IOBUF does not work->let’s debug I, O pin separately
ILA debug -> (* mark_debug = “true” * )

## Slide 68: To Fix Noise Problem
Insert attenuator, amplifier
Anlog signal -> high Q
Digital signal -> low Q
Clock length should be more shortened! (long trace -> high L -> high interference)
Digital power, analog ground segmented? -> if not, should be done
Decoupling capacitor should be placed

## Slide 69: Pin
VCC | GND | ja_3 SCLK | ja_2 IOUP2 | ja_1 SDIO2 | ja_0 CSB2
VCC | GND | ja_7 IOUP1 | ja_6 SDIO1 | ja_5 CSB1 | ja_4 PO[18]
VCC | GND | jb_3 P2[1] | jb_2 P2[0] | jb_1 PO[17] | jb_0 PO[16]
VCC | GND | jb_7 P1[2] | jb_6 P1[1] | jb_5 P1[0] | jb_4 P2[2]
GND | PROF1[2] jb_7
GND | PROF1[1] jb_6
GND | PROF1[0] jb_5
GND | IO_UPDATE1 ja_7
GND | IO_UPDATE2
GND | PROF2[0]
GND | PROF2[1]
GND | PROF2[2]
GND | RESET1
GND | SDIO1 ja_6
GND | SCLK1 ja_3
GND | CSB1 ja_5
GND | CSB2
GND | SCLK2
GND | SDIO2
GND | RESET2